APPENDIX B

VARIAN SELF ASSESSMENT CHECKLIST

NAME: ______________________________

SIGNATURE: _________________________

DATE: ______________________________

| Objectives: | Completed |
| --- | --- |
| Terminology: (2D, 2D/2D, 3D/3D, kVS, kVD, OBI, CBCT, FOV, PV) |  |
| QA: |  |
| Perform unit start up and shut down procedures |  |
| Use correct sequence for logging into and off of computers |  |
| Perform all aspects of morning QA |  |
| MLC initialization |  |
| MACHINE OPERATION IN ROOM: |  |
| Understanding of emergency procedures - Emergency buttons - Emergency Pendant/ Emergency bed procedure - Touch guards/Collision plates - LPO (Last Person Out) |  |
| Operation of the treatment couch - Understanding of all functions & buttons |  |
| Operation of Imaging Devices (Portal Vision & OBI) - Understanding of all Hand Pendant functions - Charging of the hand pendant - Imaging field sizes (e.g. 26.7 x 20 @ 150 SID) - Position of electrical wires in MV & kVD panel - Tube warm up (done in morning QA or after 4 hrs of inactivity) - Extract and retract panels - Offsetting of MV & kVD panels - Avoiding collisions - Avoiding electronics in panel - Adjusting area in interest for imaging - CBCT Modes: - Half Scan vs Full Scan - Bowtie filters (half/full) |  |
| Wedges |  |
| MACHINE OPERATION OUTSIDE ROOM: |  |
| Operation at the Console - Move gantry/collimator - Move bed - Retract imaging panels - Beam on (MV and kV) - Interrupt beam (MV and kV) - Terminate beam |  |
| Operation of the 4D ITC Console & Treatment Workflow with electronic health record |  |
| Understand the Status ball (e.g. meatball) and how to clear the faults |  |
| On Board Imager (OBI) |  |
| Acquisition of kV exposures (2D & 2D/2D) - Selecting 2D vs 2D/2D - Verify reference image is loaded on OBI workstation - Selecting appropriate fields for imaging (dragging from Dynamic Window) - Selecting Anatomy for exposure - Track changes - Download Axes - Foot switch or hand control |  |
| Acquisition of MV exposures (2D & 2D/2D) - Selecting 2D vs 2D/2D - Adding an imaging session on the 4DITC (EPI, cassette) - Verify reference image is loaded on OBI workstation - Selecting appropriate fields for imaging (dragging from Dynamic Window) - Entering MUs for exposure (2 MUs) - Entering a Delta/open exposure f/s - Manually entering field size to avoid electronics |  |
| Analyzing 2D images - Identify the reference images - Identify the MV or kV images - Adjusting image quality with Windows and Levelling (toolbar) or using Filters - Identify which is the active window - Performing a manual match - Understand viewing tools: blending, split window, moving window - Applying shifts - Deleting unwanted images (e.g. planned image only) - Save Match vs Save Image |  |
| 3D/3D CBCT Acquisition: - Understand the safe zone (e.g. Lat +/- 5, Vert -19) - Understand the importance of testing gantry clearance for the complete rotation for CBCT (in safe zone if applicable) and treatment position - Understand which filter is required with what technique - Treatment couch in ISO position before selecting 3D/3D - CBCT Wizard - Acquiring New Scan, Quick Scan (Param Display), Reconstructing an Existing Scan - New Scan: Select the CBCT Mode (e.g. Standard Dose Head) - CBCT Offset (to view anatomy outside of FOV) - Moving treatment couch in safe zone - kVD offsetting - Rotating gantry to start position - Starting and completing data acquisition - Pause data acquisition and resume - Accept and export data |  |
| Analyzing 3D/3D images: - Identify the reference & acquired image - Using the Dynamic Window - Adjust image quality with windows and filters (bone, lung) - Identify which is the active window - View slices in different planes (transverse, frontal, sagital) - Scrolling through slices & using the indicator bars to focus on a particular location - Zooming in and out - Turning contours on and off - Understand viewing tools: blending, split window, moving window - Perform Automatic Match - Setting the Clipbox/lockbox (refer to site specific imaging policies) - Select Anatomy - Intensity Range (bone or soft tissue) - Visual verification to be followed after every auto-match - Applying shifts - Save Match |  |
| Troubleshooting: |  |
| Dicom stream errors: How to clear the error |  |
| Reconstructing an existing CBCT |  |
| Reference image not appearing on OBI workstation - Reference image not available in electronic health record - Not tagged “Use for tx definition” - Icon on OBI workstation not selected |  |
| Undo an unwanted adjustment with image matching - Select Reset Shifts (returns all shifts to 0) - Control + Z (undoes the last move made) |  |
| Head Scan gantry position (22°-180°) |  |
| “Unable to perform match”message after CBCT-try resetting field edges on clip/match box & hit auto match again. This usually corrects the problem & allows the auto-match. |  |
| Mosaiq Users: - Loading reference data (Reference Planning CT) into OBI workstation from Mosaiq for CBCT only - Treatment fields (e.g. longitudinal 100, vertical 10) - Imaging set up fields - Reference CTs - Code Capturing (kV and CBCT workload) - SRO (Spatial Resolution object) Entering shifts manually if SRO failed - Off line image review instructions Troubleshooting: - Warning message from electronic health record when downloading field to 4DITC: 3D image matching is not available; - Load reference data must be performed - Reference planning CT not available (CBCT IGRT not being used for patient) - CBCT field in electronic health record incorrect (e.g. Type must be CT, couch positions don’t match treatment fields) |  |